## Slide 1
A new multichannel radiochromic film dosimetry method using Bayesian estimation techniques
XXXVII Reunión Bienal de la Real Sociedad Española de Física
Zaragoza, Spain, July 15th-19th 2019
A.D. Dominguez-Muñoz*1, H. Miras2, M.I. Gallardo1
*e-mail: adominguez18@us.es
1Universidad de Sevilla (Seville, Spain) 2Hospital Universitario Virgen Macarena (HUVM, Seville)

## Slide 2
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Outline
0
Conclusions and future work
The EBT3 radiochromic film
Calibration procedure
A new Bayesian multichannel method
Preliminary results

## Slide 3
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
The EBT3 radiochromic film
1
Estructure of GAFChromic EBT3 Dosimetry Film
Technical features
- Optimum dose range: 0.2 Gy to 10 Gy
- Near tissue equivalent
- Energy-dependence: minimal response difference from 100keV into the MV range

## Slide 4
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
The EBT3 radiochromic film
1
Estructure of GAFChromic EBT3 Dosimetry Film
Technical features
- Optimum dose range: 0.2 Gy to 10 Gy
- Near tissue equivalent
- Energy-dependence: minimal response difference from 100keV into the MV range
Spectra of an EBT3 Gafchromic film exposed to 60Co gamma-rays from the HDR-irradiator
[ P. Casolaro et al / Scientific Reports volume 9, Article number: 5307 (2019)]
Taken from
Dose
Absorbance

## Slide 5
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Calibration procedure
2
Homogeneous irradiation of film strips with known doses
Image acquisition using a flatbed colour scanner
Film analysis in the three (R,G,B) colour channels (*)
(*) Analysis made using self-developed codes with
Films irradiated with standard clinical 6 MV X-ray fields at HUVM in Seville
LOT 1

## Slide 6
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Calibration procedure
2
Homogeneous irradiation of film strips with known doses
Image acquisition using a flatbed colour scanner
Film analysis in the three (R,G,B) colour channels (*)
k ≡ colour cannel (R,G,B)
(*) Analysis made using self-developed codes with
Net Optical Density (netODk)
Films irradiated with standard clinical 6 MV X-ray fields at HUVM in Seville
LOT 1

## Slide 7
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Calibration procedure
2
Homogeneous irradiation of film strips with known doses
Image acquisition using a flatbed colour scanner
Film analysis in the three (R,G,B) colour channels (*)
k ≡ colour cannel (R,G,B)
(*) Analysis made using self-developed codes with
Net Optical Density (netODk)
Films irradiated with standard clinical 6 MV X-ray fields at HUVM in Seville
LOT 1

## Slide 8
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Calibration procedure
2
Homogeneous irradiation of film strips with known doses
Image acquisition using a flatbed colour scanner
Film analysis in the three (R,G,B) colour channels (*)
k ≡ colour cannel (R,G,B)
(*) Analysis made using self-developed codes with
Net Optical Density (netODk)
·
·
·
Films irradiated with standard clinical 6 MV X-ray fields at HUVM in Seville
LOT 1

## Slide 9
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Calibration procedure
2
k ≡ colour cannel (R,G,B)
(*) Analysis made using self-developed codes with
Films irradiated with standard clinical 6 MV X-ray fields at HUVM in Seville
LOT 1
Net Optical Density (netODk)
Homogeneous irradiation of film strips with known doses
Image acquisition using a flatbed colour scanner
Film analysis in the three (R,G,B) colour channels (*)

## Slide 10
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Calibration procedure
3
Calibration curve

## Slide 11
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Calibration procedure
3
Calibration curve
No analytical inverse!!

## Slide 12
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Calibration procedure
3
Calibration curve
No analytical inverse!!

## Slide 13
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Calibration procedure
3
Calibration curve
No analytical inverse!!

## Slide 14
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Calibration procedure
3
The lowest parameter uncertainties obtained !!
 | LOT 1 | LOT 2 | LOT 3
 | 5.28 ± 0.06 | 5.07 ± 0.11 | 5.94 ± 0.12
 | 36.8 ± 0.3 | 41.6± 0.3 | 53.8 ± 0.4
 | 2.76 | 2.79 | 3.03
 | LOT 1 | LOT 2 | LOT 3
 | 5.28 ± 0.06 | 5.07 ± 0.11 | 5.94 ± 0.12
 | 36.8 ± 0.3 | 41.6± 0.3 | 53.8 ± 0.4
 | 2.76 | 2.79 | 3.03
( + different scanner)
Red channel
Calibration curve
No analytical inverse!!

## Slide 15
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Calibration procedure
3
The lowest parameter uncertainties obtained !!
X-Rays
 | LOT 1 | LOT 2 | LOT 3
 | 5.28 ± 0.06 | 5.07 ± 0.11 | 5.94 ± 0.12
 | 36.8 ± 0.3 | 41.6± 0.3 | 53.8 ± 0.4
 | 2.76 | 2.79 | 3.03
 | LOT 1 | LOT 2 | LOT 3
 | 5.28 ± 0.06 | 5.07 ± 0.11 | 5.94 ± 0.12
 | 36.8 ± 0.3 | 41.6± 0.3 | 53.8 ± 0.4
 | 2.76 | 2.79 | 3.03
Protons:
[ M.C. Battaglia et al.
Nuclear Inst. and Methods in Physics Research B 444 (2019) 117–124]
( + different scanner)
Red channel
Calibration curve
No analytical inverse!!

## Slide 16
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
A new Bayesian multichannel method
4
Pixel dose from calibration curve
k ≡ colour cannel (R,G,B)
m ≡ pixel number
LOT 1

## Slide 17
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
A new Bayesian multichannel method
4
Pixel dose from calibration curve
k ≡ colour cannel (R,G,B)
m ≡ pixel number
LOT 1
No pixel information before irradiation !!

## Slide 18
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
A new Bayesian multichannel method
4
Pixel dose from calibration curve
k ≡ colour cannel (R,G,B)
m ≡ pixel number
LOT 1
No pixel information before irradiation !!

## Slide 19
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
A new Bayesian multichannel method
4
Pixel dose from calibration curve
k ≡ colour cannel (R,G,B)
m ≡ pixel number
LOT 1
No pixel information before irradiation !!

## Slide 20
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
A new Bayesian multichannel method
4
Pixel dose from calibration curve
k ≡ colour cannel (R,G,B)
m ≡ pixel number
LOT 1
No pixel information before irradiation !!

## Slide 21
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
A new Bayesian multichannel method
4
Pixel dose from calibration curve
k ≡ colour cannel (R,G,B)
m ≡ pixel number
LOT 1
Pixel inhomogeneities affect each channel in a different way
No pixel information before irradiation !!

## Slide 22
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
A new Bayesian multichannel method
4
Pixel dose from calibration curve
k ≡ colour cannel (R,G,B)
m ≡ pixel number

## Slide 23
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
A new Bayesian multichannel method
4
Pixel dose from calibration curve
k ≡ colour cannel (R,G,B)
m ≡ pixel number
Pixel inhomogeneities parameters
active layer
non-irradiated film

## Slide 24
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
A new Bayesian multichannel method
4
Pixel dose from calibration curve
k ≡ colour cannel (R,G,B)
m ≡ pixel number
Pixel inhomogeneities parameters
active layer
non-irradiated film
triple-cannel inferred dose

## Slide 25
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
A new Bayesian multichannel method
4
Pixel dose from calibration curve
k ≡ colour cannel (R,G,B)
m ≡ pixel number
Pixel inhomogeneities parameters
active layer
non-irradiated film
Model based on Beer-Lambert Law
triple-cannel inferred dose
Dose-independent inhomogeneities parameters!!

## Slide 26
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
A new Bayesian multichannel method
4
Pixel dose from calibration curve
k ≡ colour cannel (R,G,B)
m ≡ pixel number
Pixel inhomogeneities parameters
active layer
non-irradiated film
Model based on Beer-Lambert Law
Maximum a posteriori estimation
likelihood
priori
dose measurement variance
,
,
inhomogeneities hyperparameters
triple-cannel inferred dose
Dose-independent inhomogeneities parameters!!
Normal distribution

## Slide 27
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
A new Bayesian multichannel method
4
Pixel dose from calibration curve
k ≡ colour cannel (R,G,B)
m ≡ pixel number
Pixel inhomogeneities parameters
active layer
non-irradiated film
triple-cannel inferred dose
Model based on Beer-Lambert Law
Maximum a posteriori estimation
likelihood
priori
dose measurement variance
,
,
inhomogeneities hyperparameters
Dose-independent inhomogeneities parameters!!
Perturbation theory
Normal distribution

## Slide 28
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
A new Bayesian multichannel method
4
Pixel dose from calibration curve
k ≡ colour cannel (R,G,B)
m ≡ pixel number
Pixel inhomogeneities parameters
active layer
non-irradiated film
triple-cannel inferred dose
Model based on Beer-Lambert Law
Maximum a posteriori estimation
likelihood
priori
dose measurement variance
,
,
inhomogeneities hyperparameters
Dose-independent inhomogeneities parameters!!
Perturbation theory
Normal distribution
Weighted least squares!!
More information extracted from calibration data!!

## Slide 29
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
5

## Slide 30
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
5
X-Rays

## Slide 31
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
5
A priori distribution effect
X-Rays

## Slide 32
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
5
A priori distribution effect
X-Rays
Less variation with priori information!!

## Slide 33
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
7
Dose independence of parameters
X-Rays

## Slide 34
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
7
Dose independence of parameters
X-Rays

## Slide 35
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
7
Dose independence of parameters
X-Rays

## Slide 36
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
8
Validation: Gamma Analysis
X-Rays
Criteria: 4cGy-2mm
Gamma < 1% - 90.73%

## Slide 37
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
8
Validation: Gamma Analysis
X-Rays
Criteria: 4cGy-2mm
Gamma < 1% - 90.73%
Scanner lateral artefacts

## Slide 38
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
9
IMRT plan verification
X-Rays

## Slide 39
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
9
IMRT plan verification
X-Rays

## Slide 40
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
9
IMRT plan verification
X-Rays
Major correction in damaged pixels !!

## Slide 41
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
10
IMRT plan verification
X-Rays

## Slide 42
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
10
IMRT plan verification
X-Rays
Deep damage
Deep damage

## Slide 43
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Preliminary results
10
IMRT plan verification
X-Rays
Superficial damage
Superficial damage
Deep damage
Deep damage

## Slide 44
XXXVII Reunión Bienal de la RSEF
Zaragoza, Spain, July 15th-19th 2019
Conclusions and future work
11
Multichannel methods are necessary to obtain a unique value of the dose from the scanned image together with the calibration curve.
Our multichannel method introduces the next news:
  Use of a calibration curve with non-analytic inverse
  Weighted least squares by experimental uncertainty of measured pixel
  Dose independent inhomogeneities parameters
  Addition of priori information extracted from calibration data
Results shows a good agreement with TPS calculation.
Major correction is made in damaged zones and with strong inhomogeneities.
Comparison between inhomogeneites parameters can give us an idea about its type.
The next step is to apply this technique to low energy proton radiotherapy dosimetry, in which we expected that it plays an important role.